Metrix MTX3x52 DSO 100MS/s (VISA)

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for Metrix MTX3x52 DSO 100MS/s (VISA).

It contains functions for opening, configuring, taking measurements from, and

closing the instrument.

________________________________________________________________________________

Assumptions:

For RS-232 instrument drivers:

-  the instrument is connected to the RS-232 interface.

-  the COM port, baud rate, parity, and timeout supplied to the

initialize function must match the settings of the instrument.

For ETERNET

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Metrix MTX3x52 DSO 100MS/s (VISA)

instrument.  The functions appear in alphabetical order, with a description

of the function and its Visual Basic syntax, a description of each parameter,

and a list of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Class/Panel Name:                          Function Name:

Initialize                              mtx3x52_init

Application Functions

Example                              mtx3x52_example

Configure Functions

Vertical Setup                       mtx3x52_verticalSetup

Horizontal

Horizontal Setup                  mtx3x52_horizontalSetup

Horizontal Display                mtx3x52_horizontalDisplay

FFT                               mtx3x52_FFT

Trigger

General Trigger                   mtx3x52_setGeneralTrigger

Display

Options                           mtx3x52_setOptionDisplay

Waveform                          mtx3x52_setWaveformDisplay

Cursors                           mtx3x52_setCursors

Hardcopy Setup                       mtx3x52_setHardcopy

Action/Status Functions

Autoset                              mtx3x52_runAutoset

Run/Stop                             mtx3x52_runAcquisition

Hardcopy                             mtx3x52_runHardcopy

Data

Read Cursors                         mtx3x52_readCursors

Read Measurements                    mtx3x52_readMeasurement

Waveform

Read Channel To Array             mtx3x52_readChToArr

Read Channel To File              mtx3x52_readChToFile

Read File To Array                mtx3x52_readFileToArr

Write File To Memory              mtx3x52_writeFileToMem

Transfer

Save/Recall Setup To/From File    mtx3x52_saveSetup

File Transfer                     mtx3x52_fileTransfer

Utility Functions

Write To Instrument                  mtx3x52_writeInstrData

Read Instrument Data                 mtx3x52_readInstrData

Reset                                mtx3x52_reset

Self-Test                            mtx3x52_selfTest

Error-Query                          mtx3x52_errorQuery

Error Message                        mtx3x52_errorMessage

Revision Query                       mtx3x52_revisionQuery

Close                                   mtx3x52_close

________________________________________________________________________________

Metrix MTX3x52 DSO 100MS/s (VISA)

This instrument module provides programming support for the Metrix

MTX3x52 (VISA).  The module is divided into the following functions:

Functions/Classes:

(1) Initialize:

This function initializes the instrument and sets it to a default

configuration.

(2) Application Functions: (Class)

This class contains high-level, test and measurement routines.  These

examples call other instrument driver functions to configure, start, and

read from the instrument.

(3) Configuration Functions: (Class)

This class of functions configures the instrument by setting acquisition

and system configuration parameters.

(4) Action/Status Functions: (Class)

This class of functions begins or terminates an acquisition. It also

provides functions which allow the user to determine the current status

of the instrument.

(5) Data Functions: (Class)

This class of functions transfers data to or from the instrument.

(6) Utility Functions: (Class)

This class of functions provides lower level functions to communicate

with the instrument, and change instrument parameters.

(7) Close:

This function takes the instrument offline.

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

mtx3x52_close

Function mtx3x52_close As ViStatus (ByVal instrumentHandle As ViSession)

Purpose

This function performs the following operations:

viClose (instrSession) and viClose (rmSession).

Notes:

(1) The instrument must be reinitialized to use it again.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_errorMessage

Function mtx3x52_errorMessage As ViStatus (ByVal instrumentHandle As ViSession, ByVal errorCode As ViStatus, errorMessage As ViChar)

Purpose

This function takes the Status Code returned by the instrument driver

functions, interprets it and returns it as a user readable string.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

ByVal errorCode As ViStatus

This control accepts the Status Code returned from the instrument

driver functions.

Default Value:

0 - VI_SUCCESS

errorMessage As ViChar

This control returns the interpreted Status Code as a user readable

message string.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFF0085  Unknown status.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_errorQuery

Function mtx3x52_errorQuery As ViStatus (ByVal instrumentHandle As ViSession, errorCode As ViInt32, errorMessage As ViChar)

Purpose

This function reads an error code from the instrument's error queue.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

errorCode As ViInt32

This control returns the error code read from the instrument's error

queue.

errorMessage As ViChar

This control returns the error message string read from the

instrument's error message queue.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_example

Function mtx3x52_example As ViStatus (ByVal instrumentHandle As ViSession, ByVal source As ViInt16, ByVal time_base As ViInt16, ByVal voltageRange As ViInt16, ByVal coupling As ViInt16, waveformArray As ViReal64, number_of_samples As ViInt32, xStart As ViReal64, xIncrement As ViReal64)

Purpose

This function is an example of this driver's functions use.

The instrument is set to the single shot mode.

User can change the source channel, the timebase, the input voltage range

and the input coupling.

A trigger is then sent to the instrument.

The function waits until the acquisition is completed.

Channel is read out from the instrument.

Waveform is interpolated if needed.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal source As ViInt16

This control selects the source channel.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - CH3

4 - CH4

ByVal time_base As ViInt16

This control selects the main timebase (A) value.

Valid Values:

0  - 2ns

1  - 5ns

2  - 10ns

3  - 20ns

4  - 50ns

5  - 100ns

6  - 200ns

7  - 500ns - (Default Value)

8  - 1µs

9  - 2µs

10 - 5µs

11 - 10µs

12 - 20µs

13 - 50µs

14 - 100µs

15 - 200µs

16 - 500µs

17 - 1ms

18 - 2ms

19 - 5ms

20 - 10ms

21 - 20ms

22 - 50ms

23 - 100ms

24 - 200ms

25 - 500ms

26 - 1s

27 - 2s

28 - 5s

29 - 10s

30 - 20s

31 - 50s

ByVal voltageRange As ViInt16

This control sets vertical sense for the specified channel.

Valid Values:

0  - 2mV/div

1  - 5mV/div

2  - 10mV/div

3  - 20mV/div

4  - 50mV/div

5  - 100mV/div

6  - 200mV/div

7  - 500mV/div

8  - 1V/div

9  - 2V/div

10 - 5V/div

11 - 10V/div  (Default Value)

12 - 20V/div

13 - 50V/div

14 - 100V/div

15 - 200V/div

16 - 500V/div

17 - 1kV/div

Note : This control value must be consistent with the probe value.

See the Digiscope 2000 Operating Instructions manual for more

details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER7 (0xBFFC0907)

ByVal coupling As ViInt16

This control selects the input coupling of the selected channel.

Valid Values:

0 - AC  (Default Value)

1 - DC

2 - GND

waveformArray As ViReal64

This control specifies the name of the array in which to place the

waveform data.

Notes:

(1) The Real Array must be large enough to hold the requested points.

number_of_samples As ViInt32

This control returns the number of valid points read from the

selected channel.

xStart As ViReal64

This control returns the time of the first data point in seconds

(with respect to the trigger point).

xIncrement As ViReal64

This control returns the incrementing value in seconds for the X

axis.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Channel) out of range.

BFFC0003  Parameter 3 (Timebase) out of range.

BFFC0004  Parameter 4 (Voltage Range) out of range.

BFFC0005  Parameter 5 (Coupling) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_FFT

Function mtx3x52_FFT As ViStatus (ByVal instrumentHandle As ViSession, ByVal FFT As ViBoolean, ByVal calculate As ViInt16, ByVal scale As ViBoolean)

Purpose

This function configures the FFT mode.

The following settings are available :

- FFT ON or OFF

- Calculate

- Scale

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal FFT As ViBoolean

This control valid the FFT mode.

Valid Values:

VI_OFF (0) - OFF  (Default Value)

VI_ON  (1) - ON

ByVal calculate As ViInt16

This control selects the calculate mode.

Valid Values:

0 - Rectangular  (Default Value)

1 - Hamming

2 - Hanning

3 - Blackman

ByVal scale As ViBoolean

This control selects the Scale.

Valid Values:

VI_OFF (0) - Log scale   (Default Value)

VI_ON  (1) - inear scale

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Trigger Source) out of range.

BFFC0003  Parameter 3 (Standard) out of range.

BFFC0004  Parameter 4 (Polarity) out of range.

BFFC0005  Parameter 5 (Field) out of range.

BFFC0006  Parameter 6 (line) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC0906  Parameter 6 (Line) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_fileTransfer

Function mtx3x52_fileTransfer As ViStatus (ByVal instrumentHandle As ViSession, ByVal fileName As ViString, ByVal sourceDrive As ViInt16, ByVal overwriteFile As ViBoolean)

Purpose

This function performs a file transfer.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal fileName As ViString

This function specifies the name of the data file.

Default Value: ""

Notes:

(1) When reading a file from Computer, precede all backslashes with

an additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 15-character long max + ".???" extension.

The extention must be in capital letter.

If not an error is generated :0xBFFC0002

(3) If the file already exists on the target drive and overwrite is

not selected an error is generated : 0xBFFC090E

(4) If file does not exist on the source drive an error is generated

: 0xBFFC090D

ByVal sourceDrive As ViInt16

This control selects the source drive.

Valid Values:

0 - Computer  (Default Value)

1 - MTX3x52 Memory

ByVal overwriteFile As ViBoolean

This function specifies whether or not to write over the specified

file if it already exists on disk.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Filename) out of specifications.

BFFC0003  Parameter 3 (Source Drive) out of range.

BFFC0004  Parameter 4 (Overwrite File?) out of range.

BFFC0005  Parameter 5 (Target Drive) out of range.

BFFC0800  Error opening file.

BFFC0801  Error writing to file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC090E  File already exists.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_horizontalDisplay

Function mtx3x52_horizontalDisplay As ViStatus (ByVal instrumentHandle As ViSession, ByVal trace As ViInt16, ByVal x As ViInt16, ByVal y As ViInt16)

Purpose

This function configures the horizontal display of the instrument.

The following settings are available :

- Trace

- X (for XY mode)

- Y (for XY mode)

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal trace As ViInt16

This control selects the traces to be displayed.

Valid Values:

0 - Oscilloscope  (Default Value)

1 - XY

ByVal x As ViInt16

This control selects the X input for XY mode.

Valid Values:

1 - 1  (Default Value)

2 - 2

ByVal y As ViInt16

This control selects the Y input for XY mode.

Valid Values:

1 - 1  (Default Value)

2 - 2

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Trace) out of range.

BFFC0003  Parameter 3 (X) out of range.

BFFC0004  Parameter 4 (Y) out of range.

BFFC0005  Parameter 5 (Roll) out of range.

BFFC0006  Parameter 6 (Cumul) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_horizontalSetup

Function mtx3x52_horizontalSetup As ViStatus (ByVal instrumentHandle As ViSession, ByVal range As ViInt16, ByVal run_after_delay As ViReal64)

Purpose

This function configures the horizontal settings of the instrument.

The following settings are available :

- Pretrigger

- Range (the main timebase)

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal range As ViInt16

This control selects the main timebase  value.

Valid Values:

0  - 1ns

1  - 2ns

2  - 5ns

3  - 10ns

4  - 20ns

5  - 50ns

6  - 100ns

7  - 200ns

8  - 500ns - (Default Value)

9  - 1µs

10  - 2µs

11 - 5µs

12 - 10µs

13 - 20µs

14 - 50µs

15 - 100µs

16 - 200µs

17 - 500µs

18 - 1ms

19 - 2ms

20 - 5ms

21 - 10ms

22 - 20ms

23 - 50ms

24 - 100ms

25- 200ms

26 - 500ms

27 - 1s

28 - 2s

29 - 5s

30 - 10s

31 - 20s

32 - 50s

33 - 100s

34 - 200s

ByVal run_after_delay As ViReal64

This control selects the pretrig value ().

Default Value: 0.0

Valid Values are between -10*Range and 25*Range:

Range     Max Values    Min Values      Middle Values

200s   ->  5000          -2000           -1000

100s   ->  2500          -1000           -500

50s    ->  1250           -500           -250

20s    ->  500           -200            -100

10s    ->  250           -100            -50

5s     ->  125           -50             -25

2s     ->  50            -20             -10

1s     ->  25            -10             -5

500ms  ->  12500         -5000           -2500

200ms  ->  5000          -2000           -1000

100ms  ->  25*100        -10*100         -500

50ms   ->  25*50         -10*50          -250

20ms   ->  25*20         -10*20          -100

10ms   ->  25*10         -10*10          -50

5ms    ->  25*5          -10*5           -25

2ms    ->  25*2          -10*2           -10

1ms    ->  25*1          -10*1           -5

500us  ->  25*500        -500*10         -2500

200us  ->  25*200        -10*200         -1000

100us  ->  25*100        -10*100         -500

50us   ->  25*50         -10*50          -250

20us   ->  25*20         -10*20          -100

10us   ->  25*10         -10*10          -50

5us    ->  25*5          -10*5           -25

2us    ->  25*2          -10*2           -10

1us    ->  25*1          -10*1           -5

500ns  ->  25*500        -500*10         -2500

200ns  ->  25*200        -10*200         -1000

100ns  ->  25*100        -10*100         -500

50ns   ->  25*100        -10*100         -500

20ns   ->  25*100        -10*100         -500

10ns   ->  25*100        -10*100         -500

5ns    ->  25*100        -10*100         -500

2ns    ->  25*100        -10*100         -500

1ns    ->  25*100        -10*100         -500

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Size) out of range.

BFFC0003  Parameter 3 (Pretrig) out of range.

BFFC0004  Parameter 4 (Delay) out of range.

BFFC0005  Parameter 5 (RangeA) out of range.

BFFC0006  Parameter 6 (DelayA) out of range.

BFFC0007  Parameter 7 (RangeB) out of range.

BFFC0008  Parameter 8 (DelayB) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC0907  Parameter 7 (RangeB) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_init

Function mtx3x52_init As ViStatus (ByVal ressourceName As ViRsrc, ByVal IDQuery As ViBoolean, instrumentHandle As ViSession)

Purpose

This function performs the following initialization actions:

- Opens a session to the Default Resource Manager resource and a session

to the specified device using the interface and address specified in the

Resource_Name control.

- Performs an identification query on the Instrument.

- Resets the instrument to a known state.

- Serial Parameters are: Speed = 115200, Parity = None, Data = 8 bits,

Protocol = RTS/CTS (hard)

- Sends initialization commands to the instrument that set any necessary

programmatic variables such as Headers Off, Short Command form, and Data

Transfer Binary to the state necessary for the operation of the

instrument driver.

- Returns an Instrument Handle which is used to differentiate between

different sessions of this instrument driver.

- Each time this function is invoked a Unique Session is opened.  It is

possible to have more than one session open for the same resource.

Parameter List

ByVal ressourceName As ViRsrc

This control specifies the interface and address of the device that

is to be initialized (Instrument Descriptor). The exact grammar to be

used in this control is shown in the note below.

Default Value:  "ASRL1"

Notes:

(1) Based on the Instrument Descriptor, this operation establishes a

communication session with a device.  The grammar for the Instrument

Descriptor is shown below.  Optional parameters are shown in square

brackets ([]).

Interface   Grammar

------------------------------------------------------

Serial      ASRL<port>[::INSTR]

TCPIP       TCPIP[board]::host address::23::SOCKET

The ASRL keyword is used with Serial instruments.

The default value for optional parameters are shown below.

Optional Parameter          Default Value

-----------------------------------------

port                        1 - 16

board                       0

host address                IP address

ByVal IDQuery As ViBoolean

This control specifies if an ID Query is sent to the instrument

during the initialization procedure.

Valid Range:

VI_OFF (0) - Skip Query

VI_ON  (1) - Do Query (Default Value)

Notes:

(1) Under normal circumstances the ID Query ensures that the

instrument initialized is the type supported by this driver. However

circumstances may arise where it is undesirable to send an ID Query

to the instrument.  In those cases; set this control to "Skip Query"

and this function will initialize the selected interface, without

doing an ID Query.

instrumentHandle As ViSession

This control returns an Instrument Handle that is used in all

subsequent function calls to differentiate between different sessions

of this instrument driver.

Notes:

(1) Each time this function is invoked a Unique Session is opened.

It is possible to have more than one session open for the same

resource.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (ID Query) out of range.

BFFC0003  Parameter 3 (Reset Device) out of range.

BFFC0011  Instrument returned invalid response to ID Query

BFFC0803  Error Scanning Response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_readChToArr

Function mtx3x52_readChToArr As ViStatus (ByVal instrumentHandle As ViSession, ByVal source As ViInt16, waveformArray As ViReal64, number_of_samples As ViInt32, xStart As ViReal64, xIncrement As ViReal64)

Purpose

This function reads the specified channel from the instrument and places

it in an array.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal source As ViInt16

This control selects the source channel.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

waveformArray As ViReal64

This function specifies the name of array in which to place the

waveform data.

Notes:

(1) The Real Array must be large enough to hold the requested

points.His numberr of elements is 50000.

number_of_samples As ViInt32

This function returns the number of valid points read from the

selected channel.

xStart As ViReal64

This function returns the time of the first data point in seconds

(with respect to the trigger point).

xIncrement As ViReal64

This function returns the incrementing value in seconds for the X

axis.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Interpolation) out of range.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_readChToFile

Function mtx3x52_readChToFile As ViStatus (ByVal instrumentHandle As ViSession, ByVal source As ViInt16, ByVal drives As ViInt16, ByVal fileName As ViString, ByVal overwriteFile As ViBoolean)

Purpose

This function reads a channel from the instrument and stores it as a

file.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal source As ViInt16

This control selects the source channel.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

ByVal drives As ViInt16

This control selects the destination drive.

Valid Values:

0 - Computer

1 - MTX3x52 Memory

ByVal fileName As ViString

This function specifies the name of the wave data file.

Default Value: ""

Notes:

(1) When storing a file to Computer, precede all backslashes with an

additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 15-character long max plus ".TRC" extension.The

extention must be in capital letter.

If not, an error is generated :0xBFFC0004

(3) If the target file already exists and overwrite is not selected,

an error is generated : 0xBFFC090E

ByVal overwriteFile As ViBoolean

This function specifies whether or not to write over the specified

file if it already exists.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Source) out of range.

BFFC0003  Parameter 3 (Drive) out of range.

BFFC0004  Parameter 4 (Filename) out of specifications.

BFFC0005  Parameter 5 (Overwrite File?) out of range.

BFFC0801  Error writing to file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090C  Channel is not active.

BFFC090E  File already exists.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_readCursors

Function mtx3x52_readCursors As ViStatus (ByVal instrumentHandle As ViSession, ByVal reference As ViInt16, t1_s As ViReal64, t2_s As ViReal64, dT_s As ViReal64, tphase_s As ViReal64, v1_V As ViReal64, v2_V As ViReal64, dV_V As ViReal64, vphase_V As ViReal64)

Purpose

This function reads cursors positions.

All measured quantities are in their basic units.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal reference As ViInt16

This control selects the reference channel for the cursors positions.

Valid Values:

0 - CH1  (Default Value)

1 - CH2

t1_s As ViReal64

This control outputs the position of the vertical cursor #1.

t2_s As ViReal64

This control outputs the position of the vertical cursor #2.

dT_s As ViReal64

This control outpus the difference between vertical cursors.

tphase_s As ViReal64

This control outpus the difference between vertical cursors.

v1_V As ViReal64

This control outpus the position of the horizontal cursor #1.

v2_V As ViReal64

This control outpus the position of the horizontal cursor #2.

dV_V As ViReal64

This control outpus the difference between horizontal cursors.

vphase_V As ViReal64

This control outpus the difference between horizontal cursors.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Reference) out of range.

BFFC0003  Parameter 3 (T1) out of range.

BFFC0004  Parameter 4 (T2) out of range.

BFFC0005  Parameter 5 (dT) out of range.

BFFC0006  Parameter 6 (V1) out of range.

BFFC0007  Parameter 7 (V2) out of range.

BFFC0008  Parameter 8 (dV) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_readFileToArr

Function mtx3x52_readFileToArr As ViStatus (ByVal instrumentHandle As ViSession, ByVal drives As ViInt16, ByVal fileName As ViString, waveformArray As ViReal64, number_of_samples As ViInt32, xStart As ViReal64, xIncrement As ViReal64)

Purpose

This function reads a stored waveform file and places the data in an

array.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal drives As ViInt16

This control selects the source drive.

Valid Values:

0 - Computer

1 - MTX3x52 Memory

ByVal fileName As ViString

This function specifies the name of the wave data file.

Default Value: ""

Notes:

(1) When reading a file from Computer, precede all backslashes with

an additional backslash, e.g. if the file is in the \cvi\instr

directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 15-character long max + ".TRC" extension.

The extention must be in capital letter.

If not an error is generated :0xBFFC0003

(3) If file does not exist an error is generated : 0xBFFC090D

waveformArray As ViReal64

This function specifies the name of array in which to place the

waveform data.

Notes:

(1) The Real Array must be large enough to hold the requested

points.His numberr of elements is 50000.

number_of_samples As ViInt32

This function returns the number of valid points read from the

selected file.

xStart As ViReal64

This function returns the time of the first data point in seconds

(with respect to the trigger point).

xIncrement As ViReal64

This function returns the incrementing value in seconds for the X

axis.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Drive) out of range.

BFFC0003  Parameter 3 (Filename) out of specifications.

BFFC0004  Parameter 4 (Interpolation) out of range.

BFFC0800  Error opening file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_readInstrData

Function mtx3x52_readInstrData As ViStatus (ByVal instrumentHandle As ViSession, ByVal numberBytesToRead As ViInt32, readBuffer As ViChar, numBytesRead As ViInt32)

Purpose

This function reads data from the instrument's output buffer and returns

it to the specified variable in memory.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) If valid data is not available at the instrument's output buffer when

this function is called the instrument will hang up and the function will

not return until it times out.  If the time-out is disabled this function

will hang indefinitely and it may be necessary to reboot the computer to

break out.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal numberBytesToRead As ViInt32

The number of bytes specified by this control should be greater than

or equal to the number of bytes which are to be read from the

instrument. If the actual number of bytes to be read is greater than

the number this control specifies then multiple reads will be

required to empty the instrument's output buffer.

If the instrument's output buffer is not emptied the instrument may

return invalid data when future reads are performed.

Default Value:  50 (See NOTE 2)

Notes:

(1) If NO DATA is available at the instrument's output buffer when

this function is called the instrument will hang up and the function

will not return until it times out.  If the time-out is disabled,

this function will hang indefinitely and it may be necessary to

reboot the computer.

(2) If the number of bytes expected is greater than 50 the value of

this control may be increased. If the actual number of bytes read is

larger than the declared size of the read buffer a run-time error

will be generated.

readBuffer As ViChar

The incoming data from the instrument is placed into this variable.

Notes:

(1) Because the instrument may return both numeric and text data in

response to queries this function returns the data in string format.

(2) This function does not overwrite any old data left in the string

variable from the last time the function was called. String data in

LabWindows/CVI is terminated with an ASCII null(0x0) and string

manipulation functions will only recognize data before the ASCII

null.

(3) The declared size of the string variable must be greater than the

actual number of bytes read from the instrument if it is not a

run-time error will be generated.

numBytesRead As ViInt32

This variable contains the actual number of bytes read from the

instrument. This is the value which is returned by the read function.

Notes:

(1) If the actual number of bytes read is less than the number of

bytes specified in the Number Bytes To Read control then the output

buffer has probably been emptied. If the read function fails and the

number of bytes read is 0, the most probable cause for the failure is

there was no data available at the instrument's output buffer.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_readMeasurement

Function mtx3x52_readMeasurement As ViStatus (ByVal instrumentHandle As ViSession, ByVal measurement As ViInt16, ByVal reference As ViInt16, reading As ViReal64)

Purpose

This function forces the instrument to perform selected measurement and

returns the result.

All measured quantities are in their basic units.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal measurement As ViInt16

This control specifies the measurement function.

Valid Values:

0  - Minimum Voltage     [V]

1  - Maximum Voltage     [V]

2  - Peak to Peak        [V] (Default Value)

3  - Low Voltage         [V]

4  - High Voltage        [V]

5  - Amplitude           [V]

6  - Voltage RMS Voltage [V]

7  - Average Voltage     [V]

8  - Rise Overshoot      [%]

9  - Fall Overshoot      [%]

10 - Rise Time           [s]

11 - Fall Time           [s]

12 - Pos Pulse Width     [s]

13 - Neg Pulse Width     [s]

14 - Period               [s]

15 - Frequency           [Hz]

16 - Dutycycle           [%]

17 - N

18 _ Phase

ByVal reference As ViInt16

This control selects the reference channel for the specified

measurement.

Valid Values:

0 - CH1  (Default Value)

1 - CH2

reading As ViReal64

This control outputs the result of the measurement.

All measured quantities are output in their basic units.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Measurement) out of range.

BFFC0003  Parameter 3 (Reference) out of range.

BFFC0004  Parameter 4 (Reading) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_reset

Function mtx3x52_reset As ViStatus (ByVal instrumentHandle As ViSession)

Purpose

This function resets the instrument to a known state.

See the Digiscope 2000 Operating Instructions manual for more details.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_revisionQuery

Function mtx3x52_revisionQuery As ViStatus (ByVal instrumentHandle As ViSession, instrumentDriverRevision As ViChar, firmwareRevision As ViChar)

Purpose

This function returns the revision numbers of the instrument driver and

instrument firmware, and tells the user what firmware the driver is

compatible with, for example:

This instrument driver's Revision Number is "Rev 1.0, 04/2005, CVI 6.0"

and is compatible with Instrument Firmware Revision "1.6".

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

instrumentDriverRevision As ViChar

This control returns the Instrument Driver Software Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

firmwareRevision As ViChar

This control returns the Instrument Firmware Revision.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_runAcquisition

Function mtx3x52_runAcquisition As ViStatus (ByVal instrumentHandle As ViSession, ByVal acquisition As ViBoolean)

Purpose

This function controls the acquisition system of the instrument.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal acquisition As ViBoolean

This control valid the acquisition system of the instrument.

Valid Values:

VI_OFF (0) - Stop

VI_ON  (1) - Run   (Default Value)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Wait?) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_runAutoset

Function mtx3x52_runAutoset As ViStatus (ByVal instrumentHandle As ViSession)

Purpose

This function performs an autoset of the instrument.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0803  Error scanning response.

BFFC0900  Instrument specific error.

BFFC090F  No channel is enabled.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_runHardcopy

Function mtx3x52_runHardcopy As ViStatus (ByVal instrumentHandle As ViSession, ByVal action As ViBoolean)

Purpose

This function controls the hardcopy system of the instrument.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal action As ViBoolean

This control selects whether to run the hardcopy or to abort it.

Valid Values:

VI_OFF (0) - Abort

VI_ON  (1) - Run     (Default Value)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Action) out of range.

BFFC0803  Error scanning response

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_saveSetup

Function mtx3x52_saveSetup As ViStatus (ByVal instrumentHandle As ViSession, ByVal mode As ViBoolean, ByVal drives As ViInt16, ByVal fileName As ViString, ByVal overwriteFile As ViBoolean)

Purpose

This function save or recall set-up to or from file.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal mode As ViBoolean

This control selects between Save and Recall mode.

Valid Range:

VI_OFF (0) - Recall

VI_ON  (1) - Save  (Default Value)

ByVal drives As ViInt16

This control selects the target/source drive.

Valid Values:

0 - Computer

1 - MTX3x52 Memory  (Default Value)

ByVal fileName As ViString

This function specifies the name of the setup data file.

Default Value: ""

Notes:

(1) When reading/writing a file from/to Computer, precede all

backslashes with an additional backslash, e.g. if the file is in the

\cvi\instr directory, enter the filename as "\\cvi\\instr\\anyfile".

(2) File name must be 8-character long max + ".CFG" extension.

The extention must be in capital letter.

If not an error is generated :0xBFFC0004

(3) When saving a setup : if the target file already exists and

overwrite is not selected an error is generated : 0xBFFC090E

(4) When recalling a setup : if file does not exist an error is

generated : 0xBFFC090D

ByVal overwriteFile As ViBoolean

This function specifies whether or not to write over the specified

file if it already exists on disk.

Valid Range:

VI_OFF (0) - No  (Default Value)

VI_ON  (1) - Yes

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Drive) out of range.

BFFC0004  Parameter 4 (Filename) out of specifications.

BFFC0005  Parameter 5 (Overwrite File?) out of range.

BFFC0800  Error opening file.

BFFC0801  Error writing to file.

BFFC0803  Error Scanning Response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC090E  File already exists.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_selfTest

Function mtx3x52_selfTest As ViStatus (ByVal instrumentHandle As ViSession, selfTestResult As ViInt16, selfTestMessage As ViChar)

Purpose

This function runs the instrument's self test routine and returns the

test result(s).

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

selfTestResult As ViInt16

This control contains the value returned from the instrument self

test.  Zero means success.  For any other code, see the device's

operator's manual.

selfTestMessage As ViChar

This control contains the string returned from the self test. See the

device's operation manual for an explanation of the string's

contents.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0803  Error scanning response.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_setCursors

Function mtx3x52_setCursors As ViStatus (ByVal instrumentHandle As ViSession, ByVal cursors As ViInt16, ByVal reference As ViInt16, ByVal unattached_cursors As ViBoolean, ByVal t1_s As ViReal64, ByVal t2_s As ViReal64, ByVal tphase_s As ViReal64)

Purpose

This function sets horizontal cursors.

The following settings are available :

- Cursors

- Reference

- Unattached cursors

- T1

- T2

- TPhase

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal cursors As ViInt16

This control selects the cursors mode.

Valid Values:

0 - OFF  (Default Value)

1 - ON

2 - ON+PHASE

ByVal reference As ViInt16

This control selects the reference.

Valid Values:

0 - CH1  (Default Value)

1 - CH2

ByVal unattached_cursors As ViBoolean

This control selects the unattached cursors.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On   (Default Value)

ByVal t1_s As ViReal64

This control sets the position of the vertical cursor #1.

Valid Range:

+/- 100 * horizSense (horizSense in s/div)

ByVal t2_s As ViReal64

This control sets the position of the vertical cursor #2.

Valid Range:

+/- 100 * horizSense (horizSense in s/div)

ByVal tphase_s As ViReal64

This control sets the position of the vertical cursor #3 (phase).

Valid Range:

+/- 100 * horizSense (horizSense in s/div)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Clear Cursors) out of range.

BFFC0003  Parameter 3 (Reference) out of range.

BFFC0004  Parameter 4 (T1) out of range.

BFFC0005  Parameter 5 (T2) out of range.

BFFC0006  Parameter 6 (V1) out of range.

BFFC0007  Parameter 7 (V2) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_setGeneralTrigger

Function mtx3x52_setGeneralTrigger As ViStatus (ByVal instrumentHandle As ViSession, ByVal triggerSource As ViInt16, ByVal slope As ViBoolean, ByVal coupling As ViInt16, ByVal noiseRejection As ViBoolean, ByVal type As ViInt16, ByVal level_volts As ViReal64)

Purpose

This function configures the general trigger parameters of the

instrument.

The following settings are available :

- Trigger Source

- Slope

- Coupling

- Noise Rejection

- Type

- Level

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal triggerSource As ViInt16

This control selects the trigger source to configure.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

3 - External

4 - LINE

ByVal slope As ViBoolean

This control sets the slope of the specified channel.

Valid Values:

VI_ON  (1) - Up (Default Value)

VI_OFF (0) - Down

ByVal coupling As ViInt16

This control selects the coupling mode of the specified channel.

Valid Values:

0 - AC (Default Value)

1 - DC

2 - LF Rej

3 - HF Rej

4 _ TV-Field

5 - TV-line

ByVal noiseRejection As ViBoolean

This control sets the noise rejection of the specified channel.

Valid Values:

VI_OFF (0) - Off  (Default Value)

VI_ON  (1) - On

ByVal type As ViInt16

This control selects the type of trigger.

Valid Values:

0 - Triggered

1 - Auto (Default Value)

2 - Single

ByVal level_volts As ViReal64

This control sets level of trigger for the specified channel.

Valid Range:

-7.5 * vertSense (vertSense in Volts/div)  to  7.5 * vertSense

(vertSense in Volts/div)

Default Value: 0.0 V

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Channel) out of range.

BFFC0003  Parameter 3 (Slope) out of range.

BFFC0004  Parameter 4 (Coupling) out of range.

BFFC0005  Parameter 5 (HF Rejection) out of range.

BFFC0006  Parameter 6 (Noise Rejection) out of range.

BFFC0007  Parameter 7 (Type) out of range.

BFFC0008  Parameter 8 (Level) out of range.

BFFC0803  Error scanning response.

BFFC0809  Parameter 9 (Holdoff) out of range.

BFFC0900  Instrument specific error.

BFFC0905  Parameter 5 (HF Rejection) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_setHardcopy

Function mtx3x52_setHardcopy As ViStatus (ByVal instrumentHandle As ViSession, ByVal device As ViInt16, ByVal port As ViInt16, ByVal filename As ViString, ByVal type As ViBoolean)

Purpose

This function sets hardcopy parameters.

The following settings are available :

- Device

- Port

- Filename

- Color type

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal device As ViInt16

This control selects the device language.

Valid Values:

0  - IBM Proprinter (Default Value)

1  - EPSON ESC/P2

2  - Canon

3  - HP-PCL

4  - Seiko

5  - Postscript

6  - bitmap BMP

7  - GIF

ByVal port As ViInt16

This control selects the measurement to display.

Valid Values:

0  - RS232 (Default Value)

1  - Centronics

2  - Network

3  - File

ByVal filename As ViString

This control sets the filename where to store the hardcopy.

Default Value: ""

Note : It must be an 8-character name with no extension. Extension is

automatically added according to the device language. See the

Digiscope 2000 Operating Instructions manual for more details.

ByVal type As ViBoolean

This control sets the color of pages.

Valid Values:

VI_OFF (0) - Black & White (Default Value)

VI_ON  (1) - Color

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Device) out of range.

BFFC0003  Parameter 3 (Port) out of range.

BFFC0004  Parameter 4 (Filename) out of range.

BFFC0005  Parameter 5 (Type) out of range.

BFFC0006  Parameter 6 (# of pages) out of range.

BFFC0007  Parameter 7 (Page Number) out of range.

BFFC0008  Parameter 8 (# of pens) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC0906  Parameter 6 (# of pages) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_setOptionDisplay

Function mtx3x52_setOptionDisplay As ViStatus (ByVal instrumentHandle As ViSession, ByVal grid As ViBoolean, ByVal deviceMode As ViInt16, ByVal advanced_Mode As ViBoolean, ByVal language As ViInt16, ByVal standby As ViInt16, ByVal screen_Saver As ViInt16, ByVal brightness_ As ViInt16, ByVal contrast_ As ViInt16)

Purpose

This function configures the display options of the instrument.

The following settings are available :

- Grid

- Divice Mode

- Advanced mode

- Language

- Standby

- Screen

- Brightness

- Contrast

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal grid As ViBoolean

This control sets the grid display.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On   (Default Value)

ByVal deviceMode As ViInt16

This control selects the principal mode.

Valid Values:

0 - Scope   (Default Value)

1 - Analyser

2 - Recorder

ByVal advanced_Mode As ViBoolean

This control sets the advanced mode display.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On   (Default Value)

ByVal language As ViInt16

This control selects the language used on the instrument.

Valid Values:

0 - English

1 - French  (Default Value)

2 - German

3 - Spanish

4 - Italiano

ByVal standby As ViInt16

This control sets the standby values.

Valid Values:

0 - 30min

1 - 1H

2 - 4H

3 - 24H  (Default Value)

ByVal screen_Saver As ViInt16

This control sets the sceen saver values.

Valid Values:

0 - none

1 - 15min

2 - 30min

3 - 1H  (Default Value)

ByVal brightness_ As ViInt16

This control sets level of the brightness.

Valid values:

0 to 100%

ByVal contrast_ As ViInt16

This control sets level of the contrast.

Valid values:

0 to 100%

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Grid) out of range.

BFFC0003  Parameter 3 (Color) out of range.

BFFC0004  Parameter 4 (Screen) out of range.

BFFC0005  Parameter 5 (Menus) out of range.

BFFC0006  Parameter 6 (Language) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_setWaveformDisplay

Function mtx3x52_setWaveformDisplay As ViStatus (ByVal instrumentHandle As ViSession, ByVal minMax As ViInt16, ByVal average_rate As ViInt16, ByVal displayMode As ViInt16, ByVal repetitiveSignal As ViBoolean)

Purpose

This function configures the waveform display of the instrument.

The following settings are available :

- Min/Max acquisition

- Average rate

- Mode

- Repetitive Signal

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal minMax As ViInt16

This control Valids the Min/Max Acquisition.

Valid Values:

0 - Normal_OFF   (Default Value)

1 - Envelope_ ON

ByVal average_rate As ViInt16

This control sets the average rate.

Valid Values:

0 - No averaging (Default Value)

1 - 2

2 - 4

3 - 16

4 - 64

ByVal displayMode As ViInt16

This control selects the display mode.

Valid Values:

0 - Vector   (Default Value)

1 - Envelope

ByVal repetitiveSignal As ViBoolean

This control selects the repetitive signal.

Valid Values:

VI_OFF (0) - Off

VI_ON  (1) - On   (Default Value)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Mode) out of range.

BFFC0003  Parameter 3 (Average rate) out of range.

BFFC0004  Parameter 4 (Persistence) out of range.

BFFC0005  Parameter 5 (Interpolation) out of range.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_verticalSetup

Function mtx3x52_verticalSetup As ViStatus (ByVal instrumentHandle As ViSession, ByVal input As ViInt16, ByVal state As ViBoolean, ByVal coupling As ViInt16, ByVal probe As ViInt16, ByVal voltDiv As ViInt16, ByVal offset_volts As ViReal64, ByVal bandwidth As ViBoolean)

Purpose

This function configures the selected input channel of the instrument.

The following settings are available :

- Channel

- State

- Coupling

- Probe

- Channel inversion

- Range

- Offset

- Bandwidth limitation

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal input As ViInt16

This control selects the channel to configure.

Valid Values:

1 - CH1  (Default Value)

2 - CH2

ByVal state As ViBoolean

This control selects the state of the specified channel.

Valid Values:

VI_ON  (1) - On  (Default Value)

VI_OFF (0) - Off

ByVal coupling As ViInt16

This control selects the coupling of the specified channel.

Valid Values:

0 - AC

1 - DC  (Default Value)

2 - GND

ByVal probe As ViInt16

This control selects the probe of the specified channel.

Valid Values:

1 - x 1  (Default Value)

2 - x 10

3 - x 100

4 - x 1000

ByVal voltDiv As ViInt16

This control sets vertical sense for the specified channel.

Valid Values:

0  - 2.50mV/div

1  - 5.00mV/div

2  - 10.0mV/div

3  - 20.0mV/div

4  - 50.0mV/div

5  - 100mV/div

6  - 200mV/div

7  - 500mV/div

8  - 1V/div     (Default Value)

9  - 2V/div

10 - 5V/div

11 - 10V/div

12 - 20V/div

13 - 50V/div

14 - 100V/div

15 - 200V/div

16 - 500V/div

17 - 1kV/div

Note : This control value must be consistent with the probe value.

See the Digiscope 2000 Operating Instructions manual for more

details. If it is inconsistent an error is generated :

VI_ERROR_INCONSISTENT_PARAMETER7 (0xBFFC0907)

ByVal offset_volts As ViReal64

This control sets the signal offset of the specified channel.

Valid Ranges:

vertical sense       offset range

2.5mV/div              +/- 80mV

5mV/div              +/- 80mV

10mV/div             +/- 400mV

20mV/div             +/- 400mV

50mV/div             +/- 2V

100mV/div            +/- 2V

200mV/div            +/- 2.5V

500mV/div            +/- 12.5V

1V/div               +/- 12.5V

2V/div               +/- 64V

5V/div               +/- 64V

10V/div              +/- 64V

20V/div              +/- 640V

50V/div              +/- 640V

100V/div             +/- 640V

200V/div             +/- 6400V

500V/div             +/- 6400V

1kV/div              +/- 6400V

Default Value: 0.0

ByVal bandwidth As ViBoolean

This control selects the state of the specified channel.

Valid Values:

VI_ON  (1)   (Default Value)

VI_OFF (0)

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Input) out of range.

BFFC0003  Parameter 3 (State) out of range.

BFFC0004  Parameter 4 (Coupling) out of range.

BFFC0005  Parameter 5 (Probe) out of range.

BFFC0006  Parameter 6 (Polarity) out of range.

BFFC0007  Parameter 7 (Volts/div) out of range.

BFFC0008  Parameter 8 (Offset) out of range.

BFFC0803  Error scanning response.

BFFC0809  Parameter 9 (Bandwidth) out of range.

BFFC0900  Instrument specific error.

BFFC0907  Parameter 7 (Volts/div) is inconsistent.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_writeFileToMem

Function mtx3x52_writeFileToMem As ViStatus (ByVal instrumentHandle As ViSession, ByVal fileName As ViString, ByVal target As ViInt16)

Purpose

This function writes a wave data file (in the memory oscilloscope) to the

instrument display memory.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal fileName As ViString

This control selects the target memory.

Default Value: ""

Notes:

(1) File name must be 8-character long max + ".TRC" extension. The

extention must be in capital letter.

If not an error is generated :0xBFFC0003

(2) If file does not exist an error is generated : 0xBFFC090D

ByVal target As ViInt16

This control selects the target memory.

Valid Values:

1 - M1  (Default Value)

2 - M2

3 - M3

4 - M4

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 (Drive) out of range.

BFFC0003  Parameter 3 (Filename) out of specifications.

BFFC0004  Parameter 4 (Overwrite File?) out of range.

BFFC0800  Error opening file.

BFFC0803  Error scanning response

BFFC0900  Instrument specific error.

BFFC090D  File not found.

BFFC0910  Invalid file content.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

mtx3x52_writeInstrData

Function mtx3x52_writeInstrData As ViStatus (ByVal instrumentHandle As ViSession, ByVal writeBuffer As ViString)

Purpose

This function writes commands and queries to the instrument to modify

parameters and query device settings.

Parameter List

ByVal instrumentHandle As ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

ByVal writeBuffer As ViString

The user can use this control to send common commands and queries to

the instrument. This control can also be used to write any valid

command to the instrument.

Default Value:  "*CLS\r"

Notes:

(1) The command or query to be sent to the instrument may be a

literal enclosed in double quotes terminated by '\r' i.e. "*CLS\r" or

may be contained in a variable of type string.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.